## Sheet1
Sensors Included
Sensor | Asset | Serial Number | Time Zone | Location
Seapoint analog | Glider | SRF 4185 | UTC (GMT) | Laboratory stand
Seapoint serial | Glider | SRF 4184 | UTC (GMT) | Submerged in bucket
Turner Cylops | LRAUV | 2180551 | UTC (GMT) | Submerged in bucket
Turner Cylops | Drifter | 21102551 | Local Time (EDT) | Submerged in bucket | *Calibrated according to official Turner procedure right before test
Turner Cylops | Ship Flowthrough | 21102409 | Local Time (EDT) | Laboratory stand
Temperature (Elitech GSP-6G) | N/A | EFG212108221 | Local Time (EDT) | Submerged in bucket
Test Progression
Timepoint (UTC) | Timepoint (EDT) | Expected Concentration (ppb) | Starting Volume (mL) | Calculated Addition (mL) Day of Test | What Should Have Added (mL) | Actual Added Stock Solution (mL) | Final Volume (mL) | Actual Concentration (ppb) | Cumulative Addition | Actual Concentration Calculated Another Way As Check | Addition Measured With | Uncertainty of Measured Volume | Notes
2025-07-11 18:22 | 2025-07-11 14:22 | 0 | 6000 | 0 | 0 | 0 | 6000 | 0 | 0 | 0 | 2L Volumetric |  | * Seawater blank; some water may have been lost from adjusting the flowthrough setup (Serial Seapoint flowcap was leaking)
2025-07-11 18:27 | 2025-07-11 14:27 | 0.33 | 6000 | 0.19801 | 0.19801 | 0.2 | 6000.2 | 0.33332 | 0.2 | 0.33332 | 200 uL Pipette |  | * Because adding 0.6 mL in increments of 0.2, just recorded each addition; also, adjusted pump around 18:33 UTC, lots of bubbles/turbulence
2025-07-11 18:35 | 2025-07-11 14:35 | 0.66 | 6000.2 | 0.19801 | 0.19802 | 0.2 | 6000.4 | 0.66662 | 0.4 | 0.66662 | 200 uL Pipette
2025-07-11 18:37 | 2025-07-11 14:37 | 1 | 6000.4 | 0.20402 | 0.20403 | 0.2 | 6000.6 | 0.9999 | 0.6 | 0.9999 | 200 uL Pipette |  | * After this point, not adjusting anything in the bucket, so volume should be constant
2025-07-11 18:46 | 2025-07-11 14:46 | 5 | 6000.6 | 2.4012 | 2.4014 | 2.4 | 6003 | 4.9975 | 3 | 4.9975 | 10 mL Graduated Cylinder | 0.1 mL | * Noticed large bubble periodicaly exiting Seapoint analog exit port throughout the test from here on
2025-07-11 19:05 | 2025-07-11 15:05 | 10 | 6003 | 3.003 | 3.0045 | 3 | 6006 | 9.99 | 6 | 9.99 | 10 mL Graudated Cylinder | 0.1 mL | * From this point on, started included total volume of solution (with addition) in C1V1 calculations
2025-07-11 19:15 | 2025-07-11 15:15 | 25 | 6006 | 9.0225 | 9.0316 | 9 | 6015 | 24.9377 | 15 | 24.9377 | 10 mL Graduated Cylinder | 0.1 mL
2025-07-11 19:22 | 2025-07-11 15:22 | 50 | 6015 | 15.0752 | 15.1131 | 15 | 6030 | 49.7512 | 30 | 49.7512 | 50 mL Graduated Cylinder | 0.5 mL | * Note: all graduated cylinders and pipettes were rinsed with the stock solution before using to wash; also pipette tip was rinsed in the bucket solution when adding (200 uL)
2025-07-11 19:30 | 2025-07-11 15:30 | 100 | 6030 | 30.3015 | 30.4545 | 30 | 6060 | 99.0099 | 60 | 99.0099 | 50 mL Graduated Cylinder | 0.5 mL
2025-07-11 19:38 | 2025-07-11 15:38 | 250 | 6060 | 92.2843 | 93.2308 | 92 | 6152 | 247.0741 | 152 | 247.0741 | 100 mL Graduated Cylinder | 0.5 mL
2025-07-11 19:47 | 2025-07-11 15:47 | 400 | 6152 |  |  | 100 | 6252 | 403.071 | 252 | 403.071 | 100 mL Graduated Cylinder | 0.5 mL | * Since had to do addition for 500 ppb in two parts, just waited after first addition to add additional concentration point
2025-07-11 19:53 | 2025-07-11 15:53 | 500 | 6252 | 157.7436 | 164.5263 | 57 | 6309 | 489.7765 | 309 | 489.7765 | 100 mL Graduated Cylinder | 0.5 mL
Other Data
Stock Concentration (ppb) |  | 10000
Final Bucket Solution Volume (mL) |  | 6330 | Expected 6309, however, the 1000mL graduated cylinder used to measure final bucket solution volume had error of ~10mL
x =
a * v
s - a
x = added volume of stock solution
a = concentration of addition to total solution
v = volume of current total solution
s = concentration of stock solution
x =
v * (d - c)
s - d
x = added volume of stock solution
v = volume of current total solution
d = desired final concentration of solution
c = current concentration of solution
s = concentration of stock solution
d =
c * v + s*x
v + x
x = added volume of stock solution
v = volume of current total solution
d = desired final concentration of solution
c = current concentration of solution
s = concentration of stock solution
c =
s * x
v
c = calculated concentration of solution
x = total added volume of stock solution
v = volume of total solution
s = concentration of stock solution